| Purpose of the Goods : | The Goods are procured for the purpose to provide protection against infection. |
| --- | --- |

M : Mandatory Feature					D : Desirable Feature

| 1. | General Description |  |  |
| --- | --- | --- | --- |
| 1.1 | It shall be rectangular shaped containing three-fold pleats which open to form a cup-shaped mask covering the nose and mouth areas of the face with low breathing resistance and excellent filter characteristics. (M) |  |  |
| 2. | Material Standards |  |  |
| 2.1 | The masks shall be composed of latex free, fiberglass free and hypoallergenic material. (M) |  |  |
| 2.2 | Outer Facing: Non-woven fabric with fluid repellent property. Blue / Green / Pink / Purple / Yellow / White in colour, odorless and free of lint.(M) |  |  |
| 2.3 | Filter Medium: Melt blown high quality filter media made of polypropylene or similar materials. (M) |  |  |
| 2.4 | Inner Facing: Non-woven fabric with fluid absorbent property. Void of dyes, chemicals and inks. Soft and odorless. (D) |  |  |
| 2.5 | Nose Piece: malleable flat aluminum or polyethylene covered steel wire. (M) |  |  |
| 2.6 | Elastic Ear Loop: Knitted Polyester elastic / Polyurethane covered with Nylon & Polyester Yarn. (M) |  |  |
| 3. | Measurements: |  |  |
| 3.1 | Length : 18 cm ± 5% (M) |  |  |
| 3.2 | Width : 9 cm ± 5% (M) (Not less than 16 cm when fully extended) (M) |  |  |
| 3.3 | Number of pleats: 3 (M) |  |  |
| 3.4 | Nose Strip : Not less than 100 mm x 2 mm (M) |  |  |
| 3.5 | Elastic band: Standardized ear-loop for optimal periphery seal. Ample in length, yet soft and strong for maximum comfort. (M) Suggested inner circumference : 14.5cm (D) |  |  |
| 4. | Appearance |  |  |
| 4.1 | Free of holes, cuts, tears or any other imperfections or defects that will detract the appearance of the mask or impair its serviceability. (M) |  |  |
| 4.2 | The inner facing fabric of the mask shall be folded over the longer sides of the mask and sewed / ultrasonically welded to prevent fabric delamination. (M) |  |  |
| 4.3 | The elastic band shall be ultrasonically welded / sewed at each corner. (M) |  |  |
| 4.4 | The nose piece shall be provided to afford maximum comfort and maintaining the contours of the nose of wearer. (M) |  |  |
| 4.5 | The nose piece shall be put into the mask body completely and prevent protruding out easily. (M) |  |  |
| 4.6 | Top and bottom of the mask body shall be sewed / ultrasonically welded completely. (M) |  |  |
| 4.7 | The short sides of the folded mask shall be secured by sewed / ultrasonically welded. (M) |  |  |
| 5. | Filtration Efficiency, Air Exchange Pressure Test, Fluid Resistance & Flammability Class (Upon the Government’s request, Independent Accredited Laboratory report shall be submitted to prove the product compliance) |  |  |
| 5.1 | The mask shall have a bacterial filtration efficiency (BFE) rate of not less than 95 percent efficiency with mean particle size of 3.0 microns. (M) |  |  |
| 5.2 | The mask shall have a particulate filtration efficiency (PFE) rate of not less than 95 percent efficiency with mean particle size of 0.1 micron. (M) |  |  |
| 5.3 | The value of the Differential Pressure (Delta P) test shall not greater than 4.0 mm H2O/cm2. (M) |  |  |
| 5.4 | Synthetic Blood Penetration Resistance shall be ≥ 80 mmHg (M) |  |  |
| 5.5 | Flammability class: Class 1 or Class 2 (M) |  |  |
| 6. | Each container shall be legibly marked with the following information: |  |  |
| 6.1 | The manufacturer’s name or trade mark or recognized symbol (M) |  |  |
| 6.2 | Description of the content (M) |  |  |
| 6.3 | Lot Number (M) |  |  |
| 6.4 | Manufacturing Date (M) |  |  |
| 7. | The transport container (outer carton) shall be legibly marked with the following information: |  |  |
| 7.1 | The manufacturer’s name, trade mark or recognized symbol (M) |  |  |
| 7.2 | Description of the content (M) |  |  |
| 7.3 | Lot Number (M) |  |  |
| 7.4 | Manufacturing Date (M) |  |  |
| 7.5 | Place of origin (M) |  |  |
| 7.6 | Recommended storage conditions (D) |  |  |
| 7.7 | Weight of transport container (D) |  |  |